ASSETCORE — PHÂN TÍCH STATE MACHINE & GAP REPORT

Dựa trên Tập 3 (Data & DocType Specification) và Tập 4 (Workflow, Permission & Engine Specification)

Phiên bản: 1.0  |  Ngày: 2026-05-06  |  Tác giả phân tích: AssetCore Architecture Review

# 1. State Machine chuẩn hóa theo DocType

Mỗi bảng dưới đây liệt kê đầy đủ các transition theo format: State hiện tại → Trigger/Action → State tiếp theo, kèm Guard condition, Actor, SLA và Evidence sinh ra. Ký hiệu ⚠️ đánh dấu điểm cần review; ⛔ đánh dấu dead-end.

## 1.1  Medical Asset

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| need_registered | approve_specs | specs_approved | Need Justification approved; Spec Sheet tồn tại | HTM Manager | — | Lifecycle Event: specs_approved |
| specs_approved | approve_procurement | procurement_approved | Vendor Evaluation submitted; PO created | Procurement | — | Lifecycle Event: procurement_approved |
| procurement_approved | receive | received | Purchase Receipt submit triggers auto | Procurement / System | — | Lifecycle Event: received; Asset (core) created |
| received | open_installation | installed_pending | Installation WO created | Biomed Engineer | — | Lifecycle Event: installed_pending; WO: INSTALLATION |
| installed_pending | complete_installation | installed | Installation WO state=completed; ≥2 evidences attached | Biomed Engineer | Per procurement timeline | Lifecycle Event: installed; WO closed |
| installed | pass_initial_inspection | commissioned | Initial Inspection overall_result=pass; signed by HM | HTM Manager | — | Lifecycle Event: initial_inspection_passed; Initial Inspection approved |
| installed | fail_initial_inspection | installed_failed | Initial Inspection overall_result=fail | HTM Manager | — | Lifecycle Event: initial_inspection_failed |
| installed_failed | rework | installed_pending | Failure cause documented (≥50 ký tự) | Biomed Engineer | — | Lifecycle Event; Audit Log entry |
| commissioned | release_for_use | released_for_use | Tất cả Document Records mandatory effective; ≥1 trained operator; calibration_certificate nếu calibration_required=1; Class C/D: QMS Officer co-sign bắt buộc | HTM Manager (A/B); HM + QMS Officer co-sign (C/D) | — | Lifecycle Event: released_for_use; Notification to Department |
| released_for_use | first_use | in_use | First WO or use log submitted ⚠️ Guard chưa đủ — cần first_use_log DocType | Auto / System | — | Lifecycle Event: first_use |
| in_use | open_cm | in_repair | Failure Report submitted; severity documented | Biomed Engineer | SLA theo severity (S1: 1h HC / 2h ngoài HC) | Lifecycle Event: failure_reported; CM WO created |
| in_repair | complete_cm | released_for_use | CM WO closed; validation_result=pass; cost recorded | HTM Manager | SLA Resolve (S1: 24h, S2: 72h, S3: 7 ngày) | Lifecycle Event: repaired; Audit Log |
| in_use | place_idle | idle | Reason documented ⚠️ Guard chưa có: compliance & calibration check cụ thể | Department Head | — | Lifecycle Event: placed_idle |
| idle | return_to_use | in_use | Compliance & calibration up-to-date ⚠️ Cần field-level check cụ thể: next_calibration_due > today và legal_license_status=ok | Department Head | — | Lifecycle Event: returned_to_use |
| in_use | transfer | transferred | Asset Movement record submitted; target department xác nhận | HTM Manager | — | Lifecycle Event: transferred; Asset Movement |
| transferred | confirm_receive | in_use | Receive acknowledgement từ target Department Head | Department Head (target) | — | Lifecycle Event; Audit Log |
| in_use / out_of_service | place_out_of_service | out_of_service | ⚠️ Guard rất mờ: "Auto rule / manual" — cần định nghĩa rõ trigger: auto nếu WO CM opened, hoặc manual với reason documented | System / HTM Manager | — | Lifecycle Event; Notification BE |
| out_of_service | start_repair | in_repair | CM WO opened / Failure Report linked | Biomed Engineer | — | WO CM created; Lifecycle Event |
| out_of_service | repair_done_approved | released_for_use | ⚠️ Bypass in_repair — thiếu guard: WO CM closed + validation_result=pass bắt buộc | HTM Manager | — | Lifecycle Event |
| * (any active) | decommission | retired | Decommission Record approved; Sponsor signed | HTM Manager + Sponsor | — | Lifecycle Event: retired; Decommission Record |
| retired | dispose | disposed | Disposal Record completed | HTM Manager | — | Lifecycle Event: disposed; Disposal Record |
| retired | donate | donated | Donation agreement signed | HTM Manager | — | Lifecycle Event: donated; Donation Record |
| retired | store_long_term | stored_long_term | Storage record created + re-evaluation date set | HTM Manager | — | Lifecycle Event: stored_long_term ⛔ DEAD END — không có exit |

## 1.2  AC Work Order (Unified — 7 wo_type)

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| planned | schedule | scheduled | planned_start trong tương lai; assigned_to or assigned_team set | Biomed Engineer / HTM Manager | — | Audit Log entry |
| scheduled | start | in_progress | actual_start <= now; assigned_to present | Biomed Engineer | — | Audit Log; Lifecycle Event: pm_due / calibration_due |
| in_progress | pause | paused | Reason note bắt buộc | Biomed Engineer | — | Audit Log |
| paused | resume | in_progress | — | Biomed Engineer | — | Audit Log |
| in_progress | submit_completed | completed | ≥1 evidence; action_taken filled; actual_start + actual_end set ⚠️ Guard thiếu: wo_type-specific checks (xem bảng bên dưới) | Biomed Engineer | — | Lifecycle Event theo wo_type; Audit Log |
| completed | close | closed | validation_result=pass; cost fields filled nếu wo_type=CM và severity S1/S2 ⚠️ Guard thiếu: wo_type-specific (Calibration: cert attached; Recall: all assets processed; Installation: initial inspection submitted) | HTM Manager | — | Lifecycle Event; Asset fields updated (last_pm_date…) |
| completed | request_rework | in_progress | Reason note bắt buộc (≥ 30 ký tự) | HTM Manager | — | Audit Log |
| planned/scheduled/in_progress/paused | auto_overdue | overdue | now > sla_due ⚠️ Thiếu transition: overdue → completed (cần reason_late bắt buộc) | Auto / System (every 15 min) | Tức thì khi breach | AL-001/AL-002; Lifecycle Event; Escalation L1→L3 |
| planned/scheduled/paused | cancel | cancelled | Reason note ≥ 30 ký tự; chỉ HM; không cancel in_progress hoặc completed | HTM Manager | — | Audit Log; Lifecycle Event: wo_cancelled |

### Guard bổ sung theo wo_type (tại transition completed → closed)

Bảng dưới định nghĩa guard condition riêng cho từng wo_type, bổ sung vào base guard của AC Work Order.

| wo_type | Transition | Guard bổ sung (ngoài base) | Actor verify | Lifecycle Event sinh ra |
| --- | --- | --- | --- | --- |
| PM | in_progress → completed | ≥1 evidence; action_taken; tất cả WO Tasks completed hoặc có lý do skip; downtime_hours set | Biomed Engineer | pm_completed |
| PM | completed → closed | validation_result set (pass/fail/n_a); next PM WO được lên lịch (auto) | HTM Manager | Asset.last_pm_date updated |
| CM | in_progress → completed | ≥1 evidence; action_taken ≥ 50 ký tự; root_cause documented; downtime_hours set | Biomed Engineer | repaired |
| CM | completed → closed | validation_result=pass; cost_labor + cost_parts filled nếu S1/S2; linked_capa nếu RPN cao ⚠️ CAPA trigger rule chưa có threshold cụ thể | HTM Manager | Asset state: in_repair → released_for_use |
| CALIBRATION | in_progress → completed | ≥1 evidence; Calibration Result record linked; test_points all filled | Biomed Engineer | calibration_completed hoặc calibration_failed |
| CALIBRATION | completed → closed | Calibration Certificate attached (Document Record loại CERT-CAL); validation_result=pass; nếu regulatory: hard-stop nếu fail ⚠️ Hard-stop logic chưa định nghĩa | HTM Manager | Asset.last_calibration_date; next_calibration_due updated |
| INSPECTION | in_progress → completed | ≥1 evidence; overall_result set trong WO (ngoài Initial Inspection record) | Biomed Engineer | — (Initial Inspection DocType xử lý riêng) |
| INSPECTION | completed → closed | Inspection result đã được linked Initial Inspection record đánh dấu approved/rejected | HTM Manager | Asset state transition (commissioned / installed_failed) |
| INSTALLATION | in_progress → completed | ≥2 evidences (Tập 4); action_taken; all installation tasks done | Biomed Engineer | installed |
| INSTALLATION | completed → closed | Initial Inspection submitted (not necessarily approved) ⚠️ Tập 4 nói "Initial Inspection submitted" — chưa rõ draft hay submitted state | HTM Manager | Asset state: installed_pending → installed |
| RECALL | in_progress → completed | Tất cả affected assets trong scope được processed (WO sub-tasks per asset) ⚠️ Chưa có cơ chế track per-asset progress trong RECALL WO | Biomed Engineer | recall_initiated per asset |
| RECALL | completed → closed | All affected assets xác nhận resolved; CAPA Case linked nếu safety recall | HTM Manager | Recall Notice status updated |
| RETIREMENT | in_progress → completed | Decommission Record created; method chosen (dispose/donate/store) | HTM Manager | retired |
| RETIREMENT | completed → closed | Decommission Record signed by Sponsor; asset.state=retired | HTM Manager + Sponsor | Lifecycle Event: retired |

## 1.3  Document Record

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| draft | submit_for_review | in_review | ≥1 file attached; owner_user set ⚠️ Guard thiếu: version format valid (X.Y) | Document Owner | — | Notification to QMS Officer; Audit Log |
| in_review | approve | approved | approved_by filled; approved_by ≠ owner_user (4-eyes) ⚠️ Guard chưa enforce 4-eyes principle | QMS Officer | — | Audit Log; Approval date stamped |
| in_review | request_changes | draft | Review comments documented | QMS Officer | — | Audit Log; Notification to Owner |
| approved | set_effective | effective | effective_from date reached (auto scheduler) ⚠️ Guard thiếu: nếu effective_from < today khi approve → immediate transition cần explicit action | Auto / System (daily scheduler) | — | Lifecycle Event: document_effective; Distribution alert sent; Training triggered nếu MAJOR change |
| effective | supersede | superseded | New version of same scope transitions to effective → auto-retire old | System (on_submit new version) | — | Lifecycle Event: document_superseded; Old doc flagged |
| effective | retire | retired | Owner decision; phải có reason documented | Document Owner | — | Audit Log; Notification Distribution List |
| effective | expire | expired | effective_until < today (auto scheduler daily) ⚠️ Dead end nếu không có bản mới: expired → nên có path → retired | Auto / System | — | Lifecycle Event: document_superseded / notification AL-003/AL-004 |
| expired | supersede_expired | superseded | New version approved & effective | System | — | Lifecycle Event: document_superseded |

## 1.4  CAPA Case

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| open | approve_plan | in_action | CAPA actions ≥1; due_date set; owner_user set ⚠️ Guard thiếu: action_type bao gồm cả corrective | QMS Officer | — | Audit Log; Notification CAPA Owner |
| in_action | submit_actions_completed | awaiting_eff_check | Tất cả CAPA Actions completed_at set; evidence attached ⚠️ Actor chưa rõ: Owner hay QMS? | CAPA Owner / QMS Officer | — | Audit Log; Effectiveness Check record created |
| awaiting_eff_check | pass_effectiveness | closed | Effectiveness Check conclusion=effective; method-specific criteria met | QMS Officer | Trong 90 ngày kể từ close actions | CAPA closed; Lifecycle Event: capa_closed; Notification |
| awaiting_eff_check | fail_effectiveness | reopened | Effectiveness Check conclusion=not_effective ⚠️ Actor chưa rõ | QMS Officer | — | Audit Log; Notification CAPA Owner |
| reopened | add_new_actions | in_action | Ít nhất 1 new CAPA Action added; updated due_date set | CAPA Owner / QMS Officer | — | Audit Log |

## 1.5  Compliance Case

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| open | assign_investigate | investigating | assigned_to set; RCA technique chosen; deadline set | QMS Officer / HTM Manager | 7 ngày cho severity=high | Audit Log; Notification Assignee |
| investigating | submit_rca_done | awaiting_capa | Root Cause Analysis record linked; conclusion_md ≥ 100 ký tự ⚠️ Actor chưa rõ: ai validate RCA đã đủ? | Assignee (QMS / Biomed) | — | RCA record linked; Audit Log |
| awaiting_capa | mark_resolved | resolved | linked_capa.state = closed ⚠️ Guard thiếu: kiểm tra CAPA actual closure, không chỉ state | QMS Officer | — | Audit Log; Resolution Summary filled |
| resolved | confirm_close | closed | QMS Officer confirms; resolution_summary present ⚠️ Actor thiếu trong Tập 4 — cần explicit confirm step | QMS Officer | — | Lifecycle Event; Notification relevant parties |
| open | close_no_capa | closed | severity=low; documented justification why no CAPA needed (≥50 ký tự) ⚠️ Guard chưa enforce severity check | QMS Officer | — | Audit Log; Resolution Summary |

## 1.6  Initial Inspection

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| draft | submit | submitted | ≥1 inspection item; all items have observed values; overall_result set; ≥1 evidence ⚠️ Guard thiếu trong Tập 4 | Biomed Engineer / Inspector | — | Audit Log; Notification HM |
| submitted | approve | approved | overall_result=pass; approved_by (HM) ≠ inspector; approval_date set | HTM Manager | — | Asset state → commissioned; Lifecycle Event: initial_inspection_passed |
| submitted | reject | rejected | overall_result=fail; rejection reason documented ⚠️ Guard thiếu: rejection reason field bắt buộc | HTM Manager | — | Asset state → installed_failed; Lifecycle Event: initial_inspection_failed |
| rejected | rework | draft | Failure cause documented trong Initial Inspection ⚠️ Guard thiếu | Biomed Engineer | — | Audit Log |

## 1.7  Failure Report

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| open | triage | in_triage | Biomed Engineer assigned; acknowledged within SLA per severity ⚠️ Guard chưa có triage fields bắt buộc | Biomed Engineer | Per severity SLA response time | Audit Log; Notification BE/HM |
| in_triage | open_cm_wo | wo_created | CM Work Order created và linked ⚠️ Guard: WO.state=planned confirmed | Biomed Engineer | — | CM WO created; Lifecycle Event: failure_reported |
| wo_created | cm_closed | closed | Linked CM WO.state = closed ⚠️ Guard chưa explicit: trigger là auto hook hay manual action? | System (auto on WO close) / HTM Manager | — | Lifecycle Event: repaired; Asset state updated |
| open | close_false_alarm | closed | reason documented ⚠️ Guard thiếu: min length reason; clinical_impact=none bắt buộc cho false alarm | Biomed Engineer / HTM Manager | — | Audit Log |

## 1.8  Maintenance Plan & Calibration Plan (shared base)

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| draft | submit | submitted | ≥3 tasks (DQ-COMP-003); medical_asset set; effective_from set; frequency or interval_days set ⚠️ Guard thiếu trong workflow | Biomed Engineer | — | Audit Log; Notification HM |
| submitted | approve_activate | active | HM approves; effective_from ≤ today; no other active plan for same asset+frequency ⚠️ Guard: deactivate conflict check | HTM Manager | — | Audit Log; WO generation begins; Lifecycle Events emitted |
| active | retire | retired | Reason documented; no open WO from this plan ⚠️ Actor chưa rõ; thiếu open WO check | HTM Manager | — | Audit Log; WO generation stopped |
| active | supersede | superseded | New plan for same asset submitted and activated ⚠️ Guard: auto-triggered khi new plan active — cần explicit hook | System (auto hook) | — | Lifecycle Event: plan superseded; Old plan retired |

## 1.9  Change Control

| State hiện tại | Trigger / Action | State tiếp theo | Guard condition | Actor | SLA | Evidence sinh ra |
| --- | --- | --- | --- | --- | --- | --- |
| proposed | start_assessment | assessing | Risk assessment initiated; scope documented; change_type classified | QMS Officer | — | Audit Log; Risk Assessment record started |
| assessing | approve | approved | Risk assessment complete; approver tier per risk level: Low=QMS, Medium=PO+QMS, High=Sponsor+PO+QMS ⚠️ Guard chưa enforce tier logic in state machine | PO + QMS (or Sponsor for High) | — | Audit Log; Notification Owner to implement |
| assessing | reject | rejected | Rejection reason ≥ 50 ký tự | PO / QMS Officer | — | Audit Log; Notification Proposer |
| approved | start_implementing | implementing | Implementation plan documented; assigned_to set | Change Owner | — | Audit Log |
| implementing | submit_for_verify | verifying | Implementation done; evidence attached ⚠️ Guard thiếu: ≥1 evidence requirement | Change Owner | — | Audit Log; Notification QMS Verifier |
| verifying | verify_close | closed | Verification passed; effective_date set; evidence confirmed ⚠️ Actor chưa rõ trong state table | QMS Officer / PO | — | Audit Log; Notification all stakeholders; Change Control effective |
| verifying | return_to_implement | implementing | Issues found during verification; documented reason | QMS Officer | — | Audit Log |

# 2. Work Order Engine — Chuẩn hóa Base State

Tập 4 thiết kế một Unified Work Order Engine với 7 wo_type dùng chung base state machine. Phân tích cho thấy cách tiếp cận này là ĐÚNG và nên giữ nguyên. Sự khác biệt giữa các wo_type nên được xử lý qua guard condition và post-close hooks, không phải tách riêng state machine.

Kết luận: KEEP UNIFIED BASE — không tách state machine riêng theo wo_type. Xử lý đặc thù qua: (a) wo_type-specific guard conditions, (b) validate hooks, (c) post-close hooks, (d) child tables khi cần.

| wo_type | Base state machine | Guard bổ sung tại completed → closed | Khác biệt với base | Kết luận |
| --- | --- | --- | --- | --- |
| PM | planned → scheduled → in_progress ↔ paused → completed → closed (+ overdue, cancelled) | validation_result set; all tasks done or reason; downtime_hours | Không tách riêng. Dùng base + guard | KEEP UNIFIED |
| CM | Như trên | validation_result=pass; root_cause; cost S1/S2; capa_decision | from_failure_report bắt buộc; severity field active | KEEP UNIFIED |
| CALIBRATION | Như trên | Calibration Certificate attached; is_regulatory hard-stop nếu fail | calibration_result child record; external provider nếu external type | KEEP UNIFIED + thêm hard-stop hook |
| INSPECTION | Như trên | overall_result set; Initial Inspection DocType linked | Initial Inspection là separate DocType, không phải sub-flow của WO | KEEP UNIFIED + link guard |
| INSTALLATION | Như trên | ≥2 evidences; Initial Inspection submitted | Sau closed: auto trigger Initial Inspection draft creation | KEEP UNIFIED + post-close hook |
| RECALL | Như trên | All Recall Asset Status = completed/exempted | Cần child table Recall Asset Status (G-19). CAPA nếu safety recall. | KEEP UNIFIED + child table |
| RETIREMENT | Như trên | Decommission Record created + Sponsor signed | Sau closed: asset.state = retired auto | KEEP UNIFIED + post-close hook |

# 3. Danh sách Gap cần fix

Bảng dưới liệt kê tất cả gap phát hiện khi đối chiếu state machine với tài liệu Tập 3 và Tập 4. Màu sắc hàng: đỏ = CRITICAL, cam = HIGH, vàng = MEDIUM. Cột "Priority fix" đề xuất timeline theo Wave.

Tổng: 20 gap  |  CRITICAL: 3  |  HIGH: 9  |  MEDIUM: 8

| # | DocType / Component | Gap phát hiện | Đề xuất fix | Severity | Priority fix |
| --- | --- | --- | --- | --- | --- |
| G-01 | Medical Asset | STATE DEAD END: stored_long_term không có exit transition. Asset đưa vào kho dài hạn không thể chuyển sang retired/disposed/in_use. | Thêm 3 transition từ stored_long_term: → in_use (Return: re-evaluation approved, calibration_ok), → retired (Decommission decision), → disposed (Direct disposal from storage). Thêm guard: re_eval_date reached. | CRITICAL | Wave 1 – trước go-live |
| G-02 | Document Record | STATE DEAD END (partial): expired → không có path → retired. Nếu tài liệu hết hạn mà không có bản mới, tài liệu kẹt ở expired mãi mãi, không có cách retire chính thức. | Thêm transition: expired → retired (Actor: Document Owner; Guard: documented reason, ≥7 ngày sau expired). Giữ expired → superseded như hiện tại. | HIGH | Wave 1 |
| G-03 | Medical Asset — released_for_use → in_use | MISSING GUARD: Transition "first_use" chỉ nói "Auto on first WO/use log" nhưng không định nghĩa: (a) DocType nào là use log, (b) field nào trigger, (c) ai là actor. | Tạo DocType "Asset Use Log" hoặc dùng Frappe ToDo/Task. Định nghĩa trigger: on_submit của AC Work Order (wo_type != INSTALLATION) khi medical_asset.state = released_for_use → auto transition in_use. Actor = System, reason = first WO submitted. | HIGH | Wave 1 |
| G-04 | Medical Asset — in_use → out_of_service | MISSING GUARD: Guard là "Auto rule / manual" — quá mờ. Không định nghĩa: (a) auto rule là gì cụ thể, (b) ai được manual trigger, (c) có cần reason không. | Định nghĩa rõ: Auto = hook on Failure Report submit với clinical_impact ≠ none. Manual = HTM Manager với reason ≥ 30 ký tự. Thêm field out_of_service_reason và out_of_service_at trên Medical Asset. | HIGH | Wave 1 |
| G-05 | Medical Asset — out_of_service → released_for_use | MISSING GUARD & LOGIC: Transition bypass in_repair. Không rõ WO CM đã close chưa. Trùng lặp logic với in_repair → released_for_use. | Xóa direct transition out_of_service → released_for_use. Thay bằng: out_of_service → in_repair (khi CM WO opened) → released_for_use (khi CM WO closed). Nếu không cần repair: thêm guard out_of_service → released_for_use chỉ khi out_of_service_reason = scheduled_downtime AND HTM Manager approve. | HIGH | Wave 1 |
| G-06 | AC Work Order — overdue | MISSING TRANSITION: Không có explicit transition overdue → completed trong state machine table. Rule 4.1.10 chỉ nói "phải ghi reason chậm" nhưng không định nghĩa field. | Thêm explicit transition overdue → completed với guard bổ sung: reason_late field bắt buộc (≥50 ký tự). Thêm field reason_late trên AC Work Order. Cập nhật state machine table. | HIGH | Wave 1 |
| G-07 | AC Work Order (CM) — completed → closed | MISSING GUARD — CAPA trigger: Không có threshold cụ thể để trigger CAPA. Field linked_capa là optional không có rule khi nào bắt buộc. | Thêm business rule: nếu wo_type=CM và (severity=S1 OR root_cause ≠ "user_error") → QMS Officer phải review và quyết định CAPA. Thêm field capa_decision (required/not_required/deferred) với actor = QMS Officer. | MEDIUM | Wave 1 |
| G-08 | Document Record — draft → in_review | MISSING GUARD: Không có validation trước khi submit. Không enforce: ≥1 file attached, version format, owner ≠ creator. | Thêm validate hook: len(files) ≥ 1; version matches regex "^\d+\.\d+$"; owner_user set; title ≥ 10 ký tự. Cập nhật guard trong workflow spec. | MEDIUM | Wave 1 |
| G-09 | Document Record — in_review → approved | MISSING 4-EYES GUARD: approved_by có thể trùng owner_user. Không enforce phân tách "người viết" và "người duyệt". | Thêm validate: approved_by != owner_user. Nếu cần: approved_by phải có role QMS Officer. Ghi rõ trong spec. | MEDIUM | Wave 1 |
| G-10 | CAPA Case — in_action → awaiting_eff_check | MISSING ACTOR: Không rõ CAPA Owner hay QMS Officer xác nhận "actions completed". Dễ bị owner tự đóng mà không có QMS review. | Tách thành 2 bước: (a) CAPA Owner mark all actions done, (b) QMS Officer xác nhận → transition. Hoặc: thêm field confirmed_by_qms (Check) bắt buộc trước transition. | MEDIUM | Wave 1 |
| G-11 | Compliance Case — open → close_no_capa | MISSING GUARD: Shortcut transition chỉ có điều kiện "low severity" nhưng không enforce: (a) severity=low bắt buộc, (b) resolution documented. | Thêm guard: severity=low bắt buộc (block nếu severity=high/medium); resolution_summary ≥ 100 ký tự; justification_no_capa field bắt buộc. Actor: QMS Officer only. | MEDIUM | Wave 1 |
| G-12 | Failure Report — wo_created → closed | UNCLEAR TRIGGER: Không rõ đây là auto hook hay manual action. Nếu CM WO bị cancel, FR sẽ kẹt ở wo_created mãi. | Định nghĩa: (a) auto: hook on AC Work Order on_submit khi wo_type=CM và state=closed → FR.state=closed. (b) Edge case: nếu CM WO cancelled → FR về in_triage để BE quyết định (open new WO or close false alarm). Thêm transition wo_created → in_triage (WO cancelled). | MEDIUM | Wave 1 |
| G-13 | Calibration WO — completed → closed | MISSING HARD-STOP LOGIC: Tập 4 nói "nếu regulatory: hard-stop nếu fail" nhưng không định nghĩa: field nào check regulatory, logic hard-stop là gì, action tiếp theo là gì. | Thêm field is_regulatory (từ Calibration Plan.calibration_type). Nếu is_regulatory=1 và validation_result=fail → WO không thể close, bắt buộc mở CAPA Case và asset.state → out_of_service. Logic này phải là business rule cứng, không chỉ UI hint. | HIGH | Wave 2 |
| G-14 | Adverse Event Report | MISSING WORKFLOW: DocType tồn tại trong Tập 3 nhưng không có workflow trong Tập 4. Adverse Event là sự kiện an toàn lâm sàng nghiêm trọng, bắt buộc phải có truy vết trạng thái. | Thêm workflow: draft → submitted → acknowledged → under_review → reported_to_authority / closed. Guard: submitted cần ≥50 ký tự narrative; acknowledged: QMS ≤ 24h; reported_to_authority: nếu patient_harm. | CRITICAL | Wave 1 |
| G-15 | Recall Notice | MISSING WORKFLOW: DocType tồn tại trong Tập 3 nhưng không có workflow trong Tập 4. Recall là sự kiện FSCA, phải có SLA và trạng thái rõ ràng. | Thêm workflow: received → triaged → in_progress → completed (all assets processed) → closed. Guard: triaged ≤ 4h khi nhận; in_progress: RECALL WO created; completed: all scope_assets processed. SLA theo recall.severity. | CRITICAL | Wave 1 |
| G-16 | Service Contract | MISSING WORKFLOW: DocType tồn tại trong Tập 3 nhưng không có workflow. Service Contract là tài liệu pháp lý quan trọng, cần approval và lifecycle tracking. | Thêm workflow: draft → in_review → active → expired / terminated / renewed. Guard: active: start_date set, value > 0; expired: auto khi end_date < today; terminated: reason documented. | HIGH | Wave 1 |
| G-17 | Compliance Record | MISSING WORKFLOW cho exempted state: status có giá trị "exempted" nhưng không có workflow approval cho việc exempt. Ai được exempted? Không có audit trail. | Thêm mini-workflow cho exemption: Proposed Exemption → QMS Approved Exemption → Exempted. Hoặc: thêm trường exemption_approved_by + exemption_reason bắt buộc khi status = exempted. Không được tự chuyển status = exempted mà không có approval. | HIGH | Wave 1 |
| G-18 | Maintenance Plan / Calibration Plan — draft → submitted | MISSING GUARD: Không có guard condition nào được định nghĩa trong workflow spec cho transition này. DQ-COMP-003 yêu cầu ≥3 tasks nhưng chưa được enforce trong workflow guard. | Thêm guard: tasks ≥ 3; medical_asset set; effective_from set; no duplicate active plan for same asset+frequency (cho MP); no duplicate active plan for same asset+calibration_type (cho CP). | MEDIUM | Wave 1 |
| G-19 | Work Order Engine — Recall WO | MISSING MECHANISM: RECALL WO không có cách track tiến độ per-asset. Nếu có 50 assets bị ảnh hưởng, WO hiện tại không cho biết bao nhiêu đã xử lý xong. | Thêm child table "Recall Asset Status" vào RECALL WO: medical_asset, status (pending/in_progress/completed/exempted), completed_at, notes. Transition completed chỉ cho phép khi tất cả items trong child table = completed/exempted. | HIGH | Wave 1 |
| G-20 | Work Order Engine — base state "overdue" là warning, không phải state thật | Tập 4 nói overdue "không phải state cuối; là cảnh báo" nhưng state machine lại đặt nó như một node. Điều này dẫn đến mâu thuẫn: overdue không có defined transitions như planned/in_progress. | Giải pháp: Đổi overdue thành một flag field (is_overdue: Check) + overdue_since (Datetime) thay vì state. State machine vẫn là planned/scheduled/in_progress/paused → completed → closed. Escalation dùng is_overdue=1. Hoặc: giữ overdue như pseudo-state với explicit transitions: overdue → in_progress (resume), overdue → completed, overdue → cancelled. | MEDIUM | Wave 1 |

# 4. Tóm tắt kết quả và khuyến nghị

## 4.1 Dead-end States

Hai dead-end được phát hiện:

• G-01 — Medical Asset.stored_long_term: CRITICAL. Thiết bị vào kho dài hạn không có exit. Phải bổ sung 3 transition trước go-live.

• G-02 — Document Record.expired: HIGH. Tài liệu hết hạn không có path retire nếu không có bản mới. Phải thêm expired → retired.

## 4.2 Transition thiếu guard quan trọng nhất

Top 5 gap guard cần fix ngay (Wave 1):

1. G-04 — in_use → out_of_service: Guard "Auto rule / manual" quá mờ. Phải định nghĩa trigger cụ thể.

2. G-05 — out_of_service → released_for_use: Bypass in_repair không có WO check. Phải xóa hoặc thêm guard nghiêm ngặt.

3. G-06 — WO: overdue → completed: Transition không có trong state table. Phải thêm + bắt buộc reason_late.

4. G-03 — released_for_use → in_use: "first_use" không có DocType định nghĩa. Phải tạo trigger rõ ràng.

5. G-09 — Document Record: 4-eyes principle chưa enforce. Phải thêm approved_by ≠ owner_user.

## 4.3 DocType cần workflow nhưng chưa có

Ba DocType CRITICAL/HIGH phải có workflow trước Wave 1 go-live:

• G-14 — Adverse Event Report: Sự kiện an toàn lâm sàng. Bắt buộc có workflow và SLA acknowledged ≤ 24h.

• G-15 — Recall Notice: FSCA/Safety Recall. Bắt buộc có workflow và cơ chế track per-asset progress.

• G-16 — Service Contract: Tài liệu pháp lý. Cần approval workflow và lifecycle tracking.

• G-17 — Compliance Record (exempted): Phải có approval chain cho status=exempted.

## 4.4 Work Order Engine — Kết luận

Thiết kế Unified Work Order Engine trong Tập 4 là ĐÚNG VỀ NGUYÊN TẮC. Tất cả 7 wo_type nên dùng chung base state machine: planned → scheduled → in_progress ↔ paused → completed → closed (+ overdue, cancelled).

Tuy nhiên cần 3 cải tiến:

1. Bổ sung wo_type-specific guard table (Section 1.2 bảng 2 trong tài liệu này) vào fixtures và validation hook.

2. Giải quyết mâu thuẫn "overdue là state hay flag" (G-20): Đề xuất đổi thành flag is_overdue + overdue_since thay vì state node.

3. Thêm child table Recall Asset Status cho RECALL WO (G-19) để track per-asset progress.

## 4.5 Checklist trước go-live Wave 1

| # | Hành động | DocType | Owner | Done? |
| --- | --- | --- | --- | --- |
| 1 | Fix G-01: Thêm exit transitions từ stored_long_term | Medical Asset workflow | Dev + BA | ☐ |
| 2 | Fix G-02: Thêm expired → retired | Document Record workflow | Dev + BA | ☐ |
| 3 | Fix G-03: Định nghĩa first_use trigger | Medical Asset + AC WO hook | Dev | ☐ |
| 4 | Fix G-04: Làm rõ in_use → out_of_service guard | Medical Asset workflow | BA + Dev | ☐ |
| 5 | Fix G-05: Xóa/siết out_of_service → released_for_use | Medical Asset workflow | BA + Dev | ☐ |
| 6 | Fix G-06: Thêm overdue → completed + reason_late field | AC Work Order workflow | Dev | ☐ |
| 7 | Fix G-09: Enforce 4-eyes approval | Document Record validate hook | Dev | ☐ |
| 8 | Fix G-14: Thêm Adverse Event Report workflow | Adverse Event Report | BA + Dev | ☐ |
| 9 | Fix G-15: Thêm Recall Notice workflow + child table | Recall Notice + AC WO | BA + Dev | ☐ |
| 10 | Fix G-16: Thêm Service Contract workflow | Service Contract | BA + Dev | ☐ |
| 11 | Fix G-17: Thêm exemption approval flow | Compliance Record | Dev | ☐ |
| 12 | Fix G-18: Thêm guard MP/CP draft → submitted | MP + CP workflow | Dev | ☐ |
| 13 | Fix G-19: Thêm Recall Asset Status child table | AC Work Order (RECALL) | Dev | ☐ |
| 14 | Fix G-20: Quyết định overdue = state hay flag | AC Work Order architecture | Architect + Dev | ☐ |
| 15 | Cập nhật fixtures cho tất cả workflow changes đã fix | All workflows | Dev | ☐ |